SRL Loc Spedale 10/b 52018 Castel San Niccolò AR Tel ++39 0575 5011 Fax. ++39 0575 501200

M4000MD

DIGITAL HVAC ANALYZER

User’s Manual

M-4000MD

General Description:

An advanced measurement product, developed and produced by “MEITAV”  - gives for the first time the applications to features the following measurements capabilities

- Temperature

- Humidity

- Air Velocity

- Air Volume

- Capacity

- Dew Point

| PARAMETER | METRIC |  | ENGLISH |
| --- | --- | --- | --- |
| Temperature | C | Press  | F |
| Humidity | %RH |  |  |
| Air Velocity | Meters / Second |  |  |
| Width / Length | Centimeter |  |  |
| Air volume | Cubic meter/second | Press  | CFM |
| Power | kW | Press  | BTU/H |
| Dew Point | C |  |  |

ON button:

- Press ON button (the unit will beep and LCD highlight).

- Press ON button to switch off. (the unit will beep before to switch off)

1. Measuring Temperature

1.1 Press the ON button -

TEMP will be displayed on LCD. (fig. 1)

- Place the probe in front of unit outlet (or

any other outlet) so that the sensor’s cover

is facing the outlet.

- The LCD screen will display temperature

level in C, press  to display F, press  to display Dew

Point in C. (fig.2)

Fig.1

Sensor’s cover

Fig.2

2. Measuring Humidity Level

- Press the MODE button  - HUMIDITY% will be

displayed on LCD. (fig.3)

- Place the probe in front of units outlet

(or any other outlet) so that the sensor’s

cover is facing the outlet.

- The LCD screen will display relative

humidity level in % RH.

Sensor’s cover

Fig.3

3. Measuring Air velocity

- Press the MODE button  - VELOCITY will be displayed on LCD. (fig.4).

- Place the probe in front of units outlet  (or any other outlet) so that the sensor’s cover is facing the outlet.

- The LCD screen will display Air velocity in meters/second.

Fig.4

4. Measuring Air Volume

Fig.5

- Press the MODE button  - AIR VOLUME and ENTER WIDTH (flashing) will be displayed on LCD.(fig 5)

The M4000 will request the A\C’s outlet width.

- Using  buttons enter the appropriate measurements in cm.

- Press ENTER -

AIR VOLUME and ENTER LENGTH (flashing) will be

Fig.6

displayed on LCD. (fig 6)

The M4000 will request the A\C’s outlet length.

4.4 Using  buttons enter the appropriate measurements in

cm.

4.5 If you wish to measure round grills - step 4.3

press ENTER button more than 2 sec (long press) until

DIAMETER will be displayed on LCD. (fig.7)

4.6 Place the probe in front of the units outlet so that the

sensor’s cover is facing the outlet. Press ENTER,
The M4000 will start flashing "WAIT". "WAIT" will flash for about 20 seconds ("WAIT" flashing fig.8) put the probe on the grill and when "start average" start to flash then move the probe along the unit’s outlet so as to cover the whole area, aim 	to get an average reading from all available openings*.

Fig.7

Fig.8

- Try to cover all area in less than 10 seconds  and repeat again and again for 1 minute

- ("START AVERAGE" flashing fig.9).

- At the end of one minute the M4000 will beep signaling that measurement time ended. The LCD will display the measured Air Volume in M3/SEC (press  for CFM).(fig.10)

If you wish to redo this action. Or if you wish to measure few openings press ENTER and return steps 4.2-4.7. Measure each opening as described up to now and after the last measuring press  and LCD will display the all openings Air volume (sum).

* When moving probe along the grill’s borders (A,D) keep the probe half out of the grill.

A

B

C

D

Fig.9

Fig.10

5. Measuring Capacity

IMPORTANT:

Before you begin to measure switch with MODE to HUM mode. Put the probe in front of the inlet air  (or ambient room humidity) and wait until the measure will stable (which promise that the Temp’ measure stabilized too). Do not put it at output grill - it will take long time to return to normal room temperature and humidity conditions!

Fig.11

  - Press MODE button  - CAPACITY will be displayed on

LCD (fig.11).

(It is recommended to test at HIGH FAN SPEED)

- Align the probe in front of units inlet (or ambient room temperature) and press ENTER.

The LCD will display: HUMIDITY %, CAPACITY, Probe at RETURN AIR and "WAIT" flashing (fig.12)

- The M4000 will start measuring the unit’s inlet air parameters while displaying a flashing "WAIT". After one minute of measuring, the M4000 will beep and the "WAIT"

- will stop flashing a signal that this measurement has ended.

5.3 ENTER WIDTH flashing will be displayed on

LCD.(fig.13)

Fig.12

Fig.13

Air inlet and filter

When measuring return air the sensor box should be in the center (no need to move it during measuring)

Fig.14

Fig.15

- Using the  buttons enter the outlet’s width (cms). Press ENTER (fig.14 step 3). ENTER LENGTH will flash.

- The M4000 will request the A\C’s outlet length. Using the  buttons enter the outlet’s length (in cms) (fig.14 step 4, fig.15)

- If you wish to measure round grills - in step 5.4

press ENTER button more than 2 sec (long press) until

DIAMETER will be displayed on LCD. (fig.16)

Fig.16

- Place the probe in front of the units outlet (no need to

move it). Pay attention that the sensor’s  cover is facing

the Grill. Press ENTER - VELOCITY and a flashing "WAIT" will appear. The "WAIT" will flash for 1.5 minutes for stabilize the temperature and humidity sensors at the output air. After 1.5 minute "WAIT" will stop flashing and "START AVAREGE" will flash. (fig.17)
Move the probe along the grill.

Fig.17

IMPORTANT:

Read 4.6 (and *) before starting 5.8

- Make sure you move the probe along the units outlet so as to cover the whole area, aim to get an average reading from all available openings. At the end of this minute the M4000 will beep and the LCD screen will display the result in kW.(fig.18) Press  to see BTU/H.

Fig.18

Redo measuring outlet air:
5.9  In case you wish to redo this action (return air and the grill

size are already in memory) press

ENTER and repeat sections 5.4-5.8 above.

Measuring few outlet’s opening:

- When there are several outputs openings but the same  return air, press ENTER and repeat sections 5.4 - 5.8 for each one. In this case in section # 5.7 "WAIT" will flash for shorter period. Each time a measurement ends the LCD will display the kW for that specific outlet, if you wish to know the total  kW of all the outlets measured up to now press  .

- If you wish to go over the all process again include return air you should wait for about 10 minutes and allow the humidity sensor to return to the humidity room or help it by home ventilator. Then start from section # 5.1.